## Лист1
Item | Specification(s) | Yse/No
Testing Standards | All ASTM and ISO and other equivalent international standards. A certificate for the same should be supplied.
Vertical, dual column rigid frame (both vertical and lateral) Easy access for mounting of samples and dismantling, grips changing, furnace positioning and extensometer positioning.
The load capacity: ± 1000kN
Width of the load frame: ≥ 600 mm Depth of load frame: ≥ 550 mm
Total crosshead travel: ≥1350 mm
Stiffness of the frame: ≥ 60 kN/mm
Crosshead Speed range: 0.0005-1000 mm/min (or equivalent, providing a range of ~7 orders)
Accuracy of crosshead speed: ± 0.05% of set speed measured over full speed range or better
The moving crosshead shall be driven by precision screws with zero backlash.
Load frame should fulfil the CE conformity guidelines a per 2006/42/EG
Controller | A digital closed-loop command and feedback motion control system with a high-performance AC brushless servomotor
Limit switches to set a limit on crosshead travel
Drive system travel resolution: ≤ 0.5 nm
Drive system positioning and repetition accuracy: ±2 µm or better
In addition to computer control, an operator panel for running and stopping tests should also be provided
A failsafe emergency stop switch should be provided
Load Measurement | A single 1000 kN capacity pancake type strain gauge-based load cell for both tension and compression.
oad range: 0.4 to 100% of the capacity Force measurement should confirm to ASTM E4, ISO7500-1 Standards
Load accuracy: 0.5% to ISO 7500-1 or equivalent ASTM standard
Overload capacity without permanent zero shift: 150% of the capacity or better
oad cells should be able to resist the unnecessary parasitic influences like bending moment, torque, temperature fluctuations and humidity
Limiting bending moment: ≥ 200 Nm
Load cells should be calibrated as per international standards Supplier should be able to calibrate the load cells as per compliance to NABL norms in future
Mechanical Grips (Tensile tests) | Capacity: 1000 kN or more
Temperature range: -60 to 250 ⁰C
Maximum opening: ≥ 25 mm
Option: High temperature 20 kN mechanical grip (25 mm opening) with appropriate pull rods that can go up to 1000 oC (Material: MAR M 246/ 247 or equivalent) in place of the abovementioned pull rod-mechanical grip assembly
Software | Ability to perform tensile, compression, flexure tests Include appropriate calculation list for each type of test as per relevant international test standards (a list of standards should be provided)
Graphical user interface based, meeting all of MS Windows standards
Capable of controlling the movement of the crosshead as a function of load, stress, strain or true strain in addition to displacement control.
Data acquisition and analysis | Sampling frequency: 500 Hz or more for load, displacement, up to two optional temperature channels and up to four optional strain channels. Data rates should not be affected by the number of temperature and strain channels connected.
Real time display of graph and calculated results simultaneously.
Test control software must be able to automatically store raw data and calculated results in an ASCII file. Data should be easily exported to Excel
Capability to define correction factors such as machine compliance, slack, pretension, load and gauge length.
The ability to re-analyse past test data using different calculations must be provided.The software should have the capability to save the test method along with the start position, limit positions etc. so that the machine automatically comes to the start position for testing when the file is opened.
Power | 230 V ± 10% (1Ph), 50/60 Hz.
Safety | The testing machine should satisfy relevant safety requirements for testing systems A list of same should be provided
Automatic shutdown: There must be a provision to automatically shut off the electrical power to the machine when the specimen breaks
Training and User Manuals | Training: 5 personnel need to be trained by qualified personnel of the supplier during the time of installation and commissioning (i.e., at Kishoregonj Water Development Division)
User manual (both operation and maintenance) and certificates: 1 soft copy in a CD/DVS/USB and 1 hard copy (optional but will be highly appreciated) of the detailed user manual, complete with circuit diagrams (mechanical, electronic and electrical) and calibration certificates.
Installation files: Installation DVDs/USB for the PC and the software shall be supplied along with the equipment
Warrantee and post-supply services | Warrantee: 1 year. The supplier will be responsible for service and supply of any parts that may become faulty.
AMC (Optional): Cost of two-year AMC after warranty should be provided.
Service: The supplier should have an office or an associate (agent) in Bangladesh to provide after sales service, support and maintenance. Undertaking shall be given for provision of after sales service, software updates and spares for a minimum period of 10 years after warranty
Pre-installation preparation at Kishoregonj WD Division | Detailed pre-installation requirements and delivery period should be explicitly mentioned without any ambiguity. The quotation must clearly specify make and model of the equipment. All relevant technical literature / brochures, application notes and specifications must also be provided.
Installation and Commissioning | To be done by supplier.
Acceptance Criterion | The supplier has to demonstrate all the functions of the system according to the specifications after successful installation at Kishoregonj WD Division
Track Record of Supplier | Supplier must have installed and commissioned at least 5 same or similar capacity of machines in last three years in Bangladesh
Spares (optional) | The vendor should quote for the most common spares that will be required for 5 years of smooth operation after warranty (mention the spares, cost & use of individual items.The quoted prices must be valid for a minimum period of 3 months from the date of opening of the bid.